All staff reading PFT’s								October 3, 2017

I have asked the lab techs to include a DLco effort printout page with every PFT. This is a screen shot since we are unable to incorporate it into the report, but this is something I am working on with the manufacturer.  It is important when teaching the fellows to understand and recognize the quality of testing and efforts and determine if when the DLco is low it is due to a reduced quality effort by the patient.  The latest 2017 diffusion guidelines are uploaded into the HFHPULM site under the PFT tab and included along with 2005 guidelines. Here is what the fellows and staff should be evaluating.

- A minimum of 2 diffusion efforts are required that are within 2.0 units from each other.

  - > 5 maneuvers decrease diffusion by >3.5%

  - The acceptable efforts are averaged.

- Grade A tests requires the IVC being ≥90% of the FVC

  - If the second effort is not within 90%, but ≥85 and the alveolar volumes within 200 ml or 5% (whichever is greater) of the best effort that is >=90%, it is still a grade A quality test.

  - If 2 test Grad A tests are not within the reproducibility criteria, they are still averaged.

  - If only 1 grade a test is achieved, and the others suboptimal, the test can be graded A with comments indicating there was only 1 acceptable effort.

  - If efforts are only >=85%, they are grade B. Less that that are lower grades.

  - Test quality should be recorded on the report under the in comments (A, B, C, D & F).

  - IVC/FVC <80% or breath hold time <8 or >12 seconds should not be reported

- Understand the Gas Concentration vs Time graph.

  - Is there evidence of inspiratory or expiratory gas leaks on this graph.

  - Does identification of dead space bar by tracer gases appear accurate?

    - Tracer gas also used to calculate VA.

    - Dead space is automatically marked by software. If the technician elects to manually move it, they should report it in comments.

    - Adjustments while viewing time on x-axis causes large errors.

  - Examine adequacy of inspiratory hold (10 sec  2 sec).

  - Does the pressure tracing reflect Valsalva (’s DLco) or Mueller (’s DLco) maneuver?

  - Too slow of inspiration (> 4 sec) decreases DLco.

  - Exhalation of washout (dead space) thru discreet gas collection should not exceed 4 sec.

- Understand the Volume vs Time tracing graph.

  - 85% of IVC should be inhaled in < 4 seconds. Equations assume instantaneous filling.

  - Breath hold time should be 10  2 sec.

    - Starts at 30% of inspiratory effort time to middle of discreet gas collection.

  - Recognize normal tracings and the cause of alterations from normal.

    - If exhale is not to Residual Volume it could indicate air leak in system.

  - If expiratory volume is larger than inspiratory volume, then RV was not achieved before inspiratory effort maneuver to measure DLco.  COPD’er can take up to 12 sec to fully expire to RV.

    - A suboptimal IVC decreases diffusion – must be ≥ 90% of FVC

- Diffusion values are diurnal and can fall 2.2%/hr throughout day .

  - Also affected by smoking, hgb, ethanol, and menses (13% from 1st day until last).

Zachary Morris, MD  / Director: Pulmonary Function Lab